## Slide 1
Created with BioRender.com
Ventilators on the Fly
4 Basic Questions:
What is the mode?
What is the level of oxygen support?
Is ventilation appropriate?
How are the respiratory mechanics?
Authors: David Scudder, MD; Kaitlyn McLeod, MD; Arun Kannappan, MD
Section Editor: Trevor Steinbach, MD
Executive Editor: Yilin Zhang, MD
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation
Notes: David Scudder, MD - Resident, Departments of Medicine and Pediatrics, University of Colorado
Kaitlyn McLeod, MD - Resident, Department of Medicine, University of Colorado
Arun Kannappan, MD - Assistant Professor, Division of Pulmonary and Critical Care Medicine, University of Colorado
Trevor Steinbach, MD - Assistant Professor, Division of Pulmonary and Critical Care Medicine, University of Colorado
Yilin Zhang, MD - Assistant Professor, Department of Medicine, University of Washington

## Slide 2
VENTILATOR ORIENTATION
40
FiO2
PEEP
RR
Time
Pressure
VT
5
24
360
Time
Flow
Set by the clinician
Actual values measured by the ventilator
PEEP
ΔP
PPeak
22 cm H2O
VT
355 mL
RR
25 min-1
Min Vent
8.9 L/min
Inspiratory Flow
(Into Patient)
Expiratory Flow
(Out of Patient)
Peak Pressure
0
0
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics

## Slide 3
Pressure Control
Volume Control
PRVC
Pressure Support
30
FiO2
PEEP
RR
Time
Pressure
VT
5
24
360
Time
Flow
WHAT IS THE MODE?
PEEP and FiO2 are set in all ventilator modes to provide O2 support
Click for summary table of ventilator modes
General Approach
PPeak
25 cm H2O
VT
355 mL
RR
25 min-1
Min Vent
8.9 L/min
What you target
What you monitor
When to use
Classic square wave
Mandatory RR?
Square wave can help identify the mode
1
2
3
4
5
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics

## Slide 4
What you target
What you monitor
When to use
Classic square wave
Mandatory RR?
Pressure Control
40
FiO2
PEEP
RR
Time
Time
Pressure
Flow
Pcontrol
5
24
20
Set target
ΔP above PEEP
PPeak
25 cm H2O
VT
355 mL
RR
25 min-1
Min Vent
8.9 L/min
Square wave in pressure =
pressure mode
Pressure Control Mode
What you target
What you monitor
Patient comfort, precise airway pressure adjustment
When to use
Key outputs are tidal volume and minute ventilation
“Control mode” = mandatory RR
WHAT IS THE MODE?
Classic square wave
Mandatory RR?
Summary of ventilator modes
Volume Control
PRVC
Pressure Support
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics

## Slide 5
What you target
What you monitor
When to use
Classic square wave
Mandatory RR?
40
FiO2
PEEP
RR
Time
Pressure
VT
5
24
360
Volume Control
Time
Flow
PPeak
22 cm H2O
VT
360 mL
RR
25 min-1
Min Vent
9.0 L/min
Key output is peak pressure
Square wave in flow = volume mode
Volume Control Mode
What you target
What you monitor
ARDS (lung-protective tidal volumes)
When to use
Set target
tidal volume (VT)
“Control mode” = mandatory RR
WHAT IS THE MODE?
Classic square wave
Mandatory RR?
Summary of ventilator modes
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics
Pressure Control
PRVC
Pressure Support

## Slide 6
When to use
What you monitor
PRVC
40
FiO2
PEEP
RR
Time
Pressure
VT
5
24
360
Time
Flow
PPeak
22 cm H2O
VT
358 mL
RR
25 min-1
Min Vent
9.0 L/min
Key outputs are peak pressure and achieved tidal volume
“Control mode” = mandatory RR
Set target tidal volume (VT)
Pressure Regulated Volume Control (PRVC)
Combined mode. Clinician sets a target VT. Ventilator uses an algorithm to determine the pressure required to achieve target.
Square wave in pressure =
pressure mode
Lung protective tidal volumes with comfort of pressure support
When to use
WHAT IS THE MODE?
Summary of ventilator modes
What you monitor
Pressure Control
Volume Control
Pressure Support
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics

## Slide 7
When to use
What you monitor
What you target
Classic square wave
Mandatory RR?
Pressure Support
40
FiO2
PEEP
RR
Time
Pressure
PSupport
8
N/A
10
Time
Flow
PPeak
18 cm H2O
VT
355 mL
RR
18 min-1
Min Vent
8.9 L/min
Key outputs are tidal volume, RR, and minute ventilation
“Support mode” = no mandatory RR. Patient initiates all breaths.
Set pressure support (positive pressure on inspiration)
Negative deflection = patient-initiated breath
PSupport
Pressure Support Mode
Patient’s negative inspiratory pressure can result in blunted waves in pressure support
SBTs, patient providing adequate RR
When to use
What you monitor
WHAT IS THE MODE?
Summary of ventilator modes
What you target
Classic square wave
Mandatory RR?
Pressure Control
Volume Control
PRVC
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics

## Slide 8
VENTILATOR MODE SUMMARY
Mode | Classic square wave | Mandatory RR? | What you target | What you monitor | When to use
Pressure Control | Pressure | Yes | Inspiratory Pressure (PControl) | Tidal Volume (VT) and Minute Ventilation (VE) | Patient comfort, precise airway pressure adjustment
Volume Control | Flow | Yes | Tidal Volume (VT) | Peak Pressure and Plateau Pressure | ARDS (lung-protective tidal volumes)
Pressure Regulated Volume Control | Pressure | Yes | Tidal Volume (VT) | Peak Pressure, Plateau Pressure, Achieved Tidal Volume (VT) | Lung protective tidal volumes with comfort of pressure support
Pressure Support | Pressure | No | Pressure support (PSupport) | Tidal Volume (VT), Respiratory Rate (RR), and Minute Ventilation (VE) | SBTs, patient providing adequate RR
Mode
O2 support
Ventilation
Orientation
Cases
Mechanics
Pressure Control
Volume Control
PRVC
Pressure Support

## Slide 9
40
FiO2
PEEP
RR
Time
Pressure
VT
5
24
360
Time
Flow
Higher PEEP and FiO2 = higher O2 support
The ARDSNet PEEP ladder gives a sense of where the patient is with oxygen support.
PPeak
22 cm H2O
VT
355 mL
RR
25 min-1
Min Vent
8.9 L/min
FiO2 | PEEP
30% | 5
50% | 8
70% | 10
100% | 18-24
Room air
High O2 support
↑ PEEP, ↑ FiO2
WHAT IS THE LEVEL OF O2 SUPPORT?
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation
Notes: Full ARDSNet PEEP ladder: http://www.ardsnet.org/files/ventilator_protocol_2008-07.pdf

## Slide 10
35
FiO2
PEEP
RR
Time
Pressure
VT
5
18
360
Flow
PPeak
22 cm H2O
VT
355 mL
RR
20 min-1
Min Vent
7.1 L/min
1. Check the minute ventilation
3. Vent or Patient?
2. Is it appropriate?
IS VENTILATION APPROPRIATE?
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 11
35
FiO2
PEEP
RR
Time
Pressure
VT
5
18
360
Flow
PPeak
22 cm H2O
VT
355 mL
RR
20 min-1
Min Vent
7.1 L/min
IS VENTILATION APPROPRIATE?
Min Vent (VE) = VT x RR
7.1 L/min = 0.355 L x 20 min-1
Normal is 4-6 L/min
Often much higher when disease is present
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation
1. Check the minute ventilation
3. Vent or Patient?
2. Is it appropriate?

## Slide 12
35
FiO2
PEEP
RR
Time
Pressure
VT
5
18
360
Flow
PPeak
22 cm H2O
VT
355 mL
RR
20 min-1
Min Vent
7.1 L/min
IS VENTILATION APPROPRIATE?
pH/pCO2/pO2/HCO3
General rule is to maintain pH 7.2-7.5
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation
1. Check the minute ventilation
3. Vent or Patient?
2. Is it appropriate?

## Slide 13
no
yes
35
FiO2
PEEP
RR
Time
Pressure
VT
5
18
360
Flow
PPeak
22 cm H2O
VT
355 mL
RR
20 min-1
Min Vent
7.1 L/min
IS VENTILATION APPROPRIATE?
Is patient triggering breaths?
Adjust ventilator settings
VE = VT x RR
yes
no
Look for patient causes
Look for patient causes
CNS suppression
CNS injury
Pain
Salicylates
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation
1. Check the minute ventilation
3. Vent or Patient?
2. Is it appropriate?

## Slide 14
Volume-Targeted Modes
Pressure-Targeted Modes
Key Parameter: ΔP
Most lungs with good mechanics require a ΔP of
<10 cm H2O to achieve a
6 mL/kg tidal volume
Time
Pressure
ΔP = Peak – PEEP
Volume-Targeted Modes
Pressure-Targeted Modes
Key Parameters: VT and VE
Low VT and VE point to trouble with respiratory mechanics
What if ΔP > 10?
PPeak
28 cm H2O
VT
125 mL
RR
25 min-1
Min Vent
3.2 L/min
HOW ARE THE RESPIRATORY MECHANICS?
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 15
35
FiO2
PEEP
RR
Time
Pressure
VT
8
24
360
Time
Flow
PPeak
29 cm H2O
VT
355 mL
RR
24 min-1
Min Vent
8.5 L/min
HOW ARE THE RESPIRATORY MECHANICS?
ΔP = 29 – 8 = 21 cm H2O
In Volume-targeted modes, if ΔP > 10, perform an inspiratory hold
inspiratory hold
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 16
Compliance Issue
Resistance Issue
35
FiO2
PEEP
RR
Time
Pressure
VT
8
24
360
Time
Flow
PPeak
29 cm H2O
VT
355 mL
RR
24 min-1
Min Vent
8.5 L/min
Pplateau
Plat – PEEP > 10
Compliance Issue
Inspiratory hold
Resistance Issue
Peak – Plat > 3
Mucous plug
Kinked tube
Bronchospasm
Plat – PEEP > 10
Pneumothorax
Alveolar filling
Fibrosis
Mainstem intubation
Distended abdomen
HOW ARE THE RESPIRATORY MECHANICS?
In Volume-targeted modes, if ΔP > 10, perform an inspiratory hold
PPlateau
↑ΔP > 10
Peak – Plat > 3
Inspiratory Hold
Pplateau
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 17
What is the actual RR?
How much O2 support is the patient on?
What are the key outputs?
What is target parameter?
What is the mode?
PControl
What is the actual RR?
21 breaths per minute
How much O2 support is the patient on?
Low O2 support – patient is on minimal FiO2 and PEEP
1
2
3
21
FiO2
PEEP
RR
Time
Time
Pressure
Flow
Pcontrol
5
20
20
PPeak
25 cm H2O
VT
345 mL
RR
21 min-1
Min Vent
7.2 L/min
What are the key outputs?
VT and minute ventilation
Case 1
What is target parameter?
Pressure control
What is the mode?
Square wave in pressure = pressure mode
1
Mandatory RR = control mode
2
3
4
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 18
What is the actual RR?
How much O2 support is the patient on?
What are the key outputs?
What is target parameter?
What is the mode?
1
2
3
100
FiO2
PEEP
RR
Time
Pressure
VT
18
24
360
PPeak
27 cm H2O
VT
355 mL
RR
25 min-1
Min Vent
8.9 L/min
Time
Flow
Peak pressure (ΔP <10, so inspiratory hold not needed)
Square wave in flow = volume mode
1
2
VT
What is the actual RR?
25 breaths per minute
How much O2 support is the patient on?
Very high O2 support. Maximum FiO2 and PEEP.
What are the key outputs?
What is target parameter?
Volume control
What is the mode?
3
ΔP = Peak – PEEP
4
Case 2
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 19
What is the actual RR?
How much O2 support is the patient on?
What are the key outputs?
What is target parameter?
What is the mode?
2
3
1
30
FiO2
PEEP
RR
Time
Pressure
PSupport
5
N/A
10
PPeak
28 cm H2O
VT
355 mL
RR
25 min-1
Min Vent
8.9 L/min
Time
Flow
Negative deflection = patient-initiated breath
VT, RR, and minute ventilation
PSupport
What is the actual RR?
25 breaths per minute
How much O2 support is the patient on?
Low O2 support. Low FiO2 and PEEP
What are the key outputs?
What is target parameter?
Pressure support
What is the mode?
1
No mandatory RR = support mode
2
Square wave in pressure = pressure mode. Blunted waves are from patient’s negative inspiratory pressure
4
3
Case 3
Mode
O2 support
Ventilation
Cases
Mechanics
Orientation

## Slide 20
Modern Framing — Lung-Protective Ventilation (unchanged core, 2015–24 evidence)
The original talk teaches mode identification; this slide adds current ARDS ventilation targets. The standing teaching is unchanged.
Low tidal volume (still the foundation)
• 6 mL/kg predicted body weight (range 4–8) in ARDS; the original ARDSNet mortality benefit (NEJM 2000) remains standard of care.
Pressure limits
• Keep plateau pressure <30 cm H2O (many target ≤28).
Driving pressure (ΔP = Plateau − PEEP) — the newer prognostic lever
• Lower driving pressure is the ventilator variable most strongly associated with survival in ARDS (Amato, NEJM 2015; mediation analysis of 3,562 patients). Target ΔP ≤14–15 cm H2O where achievable.
• Note: on this slide ΔP = Peak − PEEP (a quick bedside screen). True driving pressure uses the PLATEAU pressure (inspiratory hold), not peak.
Sources: ARDSNet, NEJM 2000;342:1301. Amato MBP et al, NEJM 2015;372:747. Lung-protective targets unchanged through 2024.

## Slide 21
Take-Home Points
• Read any ventilator with 4 questions: (1) What mode? (2) What O2 support? (3) Is ventilation appropriate? (4) How are the mechanics?
• Mode clues: square wave in PRESSURE = pressure mode (PC, PRVC, PS); square wave in FLOW = classic volume control. A set/mandatory RR = control mode; no mandatory RR = support mode (PS).
• Target vs monitor: set pressure → monitor VT & minute ventilation; set volume → monitor peak & plateau pressure. PEEP and FiO2 set the O2 support in every mode (ARDSNet PEEP/FiO2 ladder).
• Ventilation: VE = VT × RR (normal 4–6 L/min, often higher in disease); aim to keep pH ~7.2–7.5. Check whether breaths are vent- or patient-triggered.
• Mechanics: in volume-targeted modes, if ΔP > 10 do an inspiratory hold. Peak − Plateau > 3–5 = RESISTANCE problem (mucus plug, kinked tube, bronchospasm); elevated Plateau − PEEP (>10) = COMPLIANCE problem (pneumothorax, edema/alveolar filling, fibrosis, mainstem intubation, distended abdomen).

## Slide 22
Attribution & Changelog
Original
• "Ventilators on the Fly" — Scudder D, McLeod K, Kannappan A; sec. ed. Steinbach T; exec. ed. Zhang Y. TeachIM.org, Dec 2021. https://teachim.org/teaching_material/ventilators-on-the-fly/ — CC BY-NC 4.0. Educational use only; not medical advice.
Correction made to this copy (June 2026)
• Slide 16: "Compliance Issue" box label "Plat – Peak > 10" → "Plat – PEEP > 10" — fixes an internal inconsistency with the same slide’s waveform annotation; compliance is flagged by elevated plateau relative to PEEP.
Appended (originals untouched)
• Lung-protective / driving-pressure framing slide; take-home points; this attribution slide. All interactive animations and click-to-reveal builds preserved.
Companion files
• Learner Handout (PDF) and Teachers Guide (DOCX) copied unchanged into the lesson folder & linked — reviewed, no factual errors.